Grooming A Newly Installed Q330 at a TA Station Using Willard
20090120  aws  Ver 1.1

20090617 aws Ver 1.2  - Authorization code removed for public access.

1) Check emails with the Subject line: Online Report -  xxxx, or Subject line: Service Report – xxxx.  For all stations (xxxx) that are newly online, and those stations that were serviced so that a new Q330 is on site, perform the following grooming operations to ensure that the Tag, Config, GPS, and other settings are correct and that data is flowing correctly to modem and baler, and that the sensor is centered and acting like a seismometer.

2) Connect your PC to the internet and make sure you can get to various web pages.

3) Open mail and find the Online or Service Report for the station you want to groom.  The online report has the two things you need on the first couple of lines:

      Q330 SN - 010000044DA95891
      Cell IP - 166.241.200.153

      The service report might only list the new Q330 SN.  If you don’t already have it,  

      find the serviced station IP from steps 4 thru 11.  If you know the Q330 SN and the station IP, go directly to step 16.
4) Open Willard and click on Manager: [image: image1.png]Consele Registation | D, select  stationfrom the st below

o [IEH

 Register

Descipion |

il cose |

Help ‘

18 Monager

£ unipal





5) Click on ConfigServer:

     [image: image2.png]Willard Station Manager
Select Statian to Modiy

Ditectory [CABACKUPAUSARRAY\D330NEWAWT
Station Description

Wose | ooke | wied | _cm

A Do Conligserver





6) Click on Choose:

    [image: image3.png]ConfigServer

Curent ConfigServe

o #dd or Update Station Show Station List





7) Enter http://ta-config.isti.com:8133  or  http://anfpublic.ucsd.edu:8133  into the Select Config Server and click OK.  The two servers have slightly different lists of stations – ta-config keeps some of the older, transferred stations, while anfpublic has some of the test stations at PASSCAL.  
    [image: image4.png]Select ConfigServer

User Name

Passward

hdd Updale Dekete

X Cancel





8) Click on Show Station List:

    [image: image5.png]ConfigServer
Curent ConfgServer Commuricaions
g/ aconfgisti com 133

Show Station List




9) A list of TA stations should appear.  Click on the button above DP4 to get the list ordered:

  [image: image6.png]ConfigServer

Curent ConfigServer Commurications

hitp:/taconfigisticom 8133 Choase || Done

Stations
Sort Rlealm Requested

& €| & Auomatic C Local C Intemet
TAG D DP 4 [Desciipiion A
o0 TAKIEA
1005 TaHAST
1006 AL
1009 AL
1044 TaZ1A
1102 TADEA
1102 TASUTE
1103 TaFI0s
1105 T2
1106 TAELFS
1107 TAHATC
1109 TARTIA
1111 TAHELL
1118 TaS08C
110 AN

Show Station List





8)  Scroll down until the station you wish to connect to appears and click on the station name under DP4:

[image: image7.png]ConfigServer

Curent ConfigServer Commurications

hitp:/taconfigisticom 8133 Choase || Done

Stations
Sort Rlealm Requested

» & | & Auomaic C Local C Intemet
4G D DP 4 [Desciipiion A
=0 Ta-COEA
1886 TALO7A
1914 TAC0GA
1608 THCIGH,

1834 TACI0A
1591 TALIZA
1912 UEY
1770 TAC1A
1987 TAD03
1140 D04
1575 D05
1902 TAD0BA
1908 TAD07A
1968 D08
1 TAnNR

Show Station List





10) Click on where it says Ethernet:

[image: image8.png]ConfigServer

Curent ConfigServer Commurications

hitp:/taconfigisticom 8133 Choase || Done

Station Detal

= Statin
Latiude=47.8007
Longiude=118.2741
Elevalion=0.4848
=03
Tog=1608
Seal Number=0100000B637B3764

Base Por-5330
Web Port-5380

P2
oP3
= DP4
Type=Balerl4
ID=TA LD
1P Address=70.41.247.239
Web Port-5354

/' Add or Update Station Show Station List





11) Click on “Add or Update Station” above to get the following window:

[image: image9.png]Willard Station Manager
St Saon ooty
10 595 =] Discto |

Station Description

Goss | odse | s | |

X Coa





12) Click Create above.  Note: If you’ve already created this station in your Willard list, you will not get this box, but will jump to the next box:

[image: image10.png]Q330 Configuration for. DOGA
1P Adkdresses or Damain Names 0330 Inteface to Use
seidt || C seialPatl
seid2z || C sePo2
Ethemet [166.16198244 | Ethemet

0330 Seril Number [OTOD00IDB1G4ESS2  BaselP Pot [5320

7 =





13) Click OK:

[image: image11.png]Willard Configuration for DOGA

Curent Intetface
 Diect Seril

 Winsock (etwork) £ bodem

Seial Inerface Dptions
RTS [ CTS [ iwed CTSTmef0 192 (2]

General Dptions
Seil P [10172535 Command Timeout
64 104

1P adeh of PC's COM part Min x| Man v

Time and Date Formating
Date [ypymmdd Time [Htrrrss

Configuration and Token Fil Format
© & Verboss XML Old Stle CFG/TOK

7





14) Click OK again.  Note: you will not get the above box, if you’ve already created this station in your Willard list, but will jump to the next box:

         [image: image12.png]Willard Station Manager
Select Statian to Modiy

o [oos

Diectoy [CABACKUPUSARFAT TONEWWI

Station Description

e |

oo | wied | o

A Do Conligserver





15) Click Close above
[image: image13.png]Consele Registation | D, select  stationfrom the st below

o [T

 Register

Descipion |

il cose |

Help ‘

18 Monager

£ unipal





16) Verify that you have the up-to-date IP and Q330 SN by clicking on “Manager” and enter in the ID of the station you are about to groom(in this example V29A) [image: image14.png]Willard Station Manager
Select Statian to Modiy

Ditectory [CABACKUP\USARRAY\D330WILLA
Station Description

Wose | ooke | wied | _cm

A Do Conligserver





17) Click on the Q330 button to see: [image: image15.png]Q330 Configuration, for, V29A
Standerd IP Addresses or Domain Names 0330 Inteface to Use:
seid1 [0000 || SeialPotl
seid2 0000 || O SeilPot2
Ethemet [166.241200153 | Ethemet

0330 Seril Number [010000044DASS3ST  Standard Base Part [5320

Custom Adcress:Port Paits:

E— ] Y

v





18) Make sure the Ethernet address matches the most current IP address and that the Q330 Serial number is correct.  If the new service report shows a different SN for the new Q330, type the new one in.

19) Click ok once the settings are correct, then click Close to leave the manager window.

 [image: image16.png]hoose Station.

Consele Registation | D, select  stationfrom the st below

o [z

' Fegita |

Descipion |

FLoese |

Help ‘

B e |

£ Unipal |





20) Use the arrow and click on the station you want to connect to in the ID list.  Click Register

[image: image17.png]V29A Registration to 166.241.200.153

Q330Pot
& Configuation

€ Special Functions

I Use P Broadcast
Wilard Inteface on the PC.

& Winsock [network)
I~ Tunnel330 (TCP)

AT

0330 Inteface
@ Ethemet 166.241.200.153
4
o

o =

Authorizatin Code

¥ Save Autharizalion Cads.

@ ooppg | Mg |





21) Enter the Authorization Code (ask a TA administrator for it), and click on Save Authorization code.  Click Register.  

22) You should now connect to the remote Q330 station. [image: image18.png]Willard Q330 Setup & Control - Station V29A (166.241.200.153)
Commands  Information  Status  Configuration  Help

CaL Status ) Phose [ s Cock [IN]

Total Packets 15 Checksum Enors
Crd Timeauts 0 UDEnore

Out of Window 0 Average Received Bps
2009.01-20 2351:39.000000

20090120 16:57:32 Interface Progranming Feseived

2009.01-20 16:51:34.Thread Names ecelved

2009.01-20 16:57: 5 Manufacturer's Data Received, Propety Tag=718
2009.01-20 16:51:36 Fhed Values Received

2009.01-20 16:51:37.Globa Pragramming Received

2009.01-20 165738 Data Port 1 Programming Received

2009.01-20 16:57:38Data Port 2 Programming Received

2009.01-20 16:57:38Data Port 3 Programming Received

2009.01-20 16:57:38Data Port 4 Programming Received

2009.01-20 16:51:33.6PS Receiver D Received





23) Note that the Clock is at 100%.  If this isn’t so, note it in the notes you are keeping on the grooming.  I keep a one line per station file showing Station name, initials of installer/servicer, Date and current status:
      R27A bp 20081210, 1215, 20081216  fine

      S26A bp 20081210  20081215 fine 

      N26A ko 20081210 New IE fine  The reason there are several different dates for R27A and S26A is that I could not connect to the station on those earlier dates.   New IE indicates the Vault Interface Enclosure has gone in that site.

24) A station that doesn’t connect will have a Willard window that looks like the following: [image: image19.png]Willard Q330 Setup & Control - Station 1106
Commands  Information  Status  Configuration  Help

s P oo oo

Status
Total Packets 0 ChecksumEnors
Crd Timeauts 2 UDEnoe
Out of Window 0 Average Received Bps

1/20/2009 50301 PMFIF fiter DECTD loaded vih 400 pais.
1/20/2009 50301 PHM.FIR iter VL3833 loaded with 389 poirs.
1/20/2009 50301 PMFIR fiter LLP373 loaded ith 373 poirts.
1/20/2009 50301 PM.FIR fiter U520 loaded with 200 pais
1/20/2009 50301 PM.FIR it FLIS20 loaded with 32 pais
1/20/2009 50301 PM.FIR fiter FUSB0 loaded with 22 pais.




  The command Timeouts will increase to 22 and Willard will give up on that station.  Non-connection can be caused by you entering in the incorrect or outdated IP or Q330 S/N, so always check that they are correctly entered.  Non connection is most often caused by a communication breakdown between the Q330 and the internet.  If this is the case, just note the date in your grooming notes and indicate no connection was made.  If this persists beyond a few days after install or servicing, it needs to be reported so a visit can be made to determine the problem.
25) If Willard appears more like the screenshot in step 22, the station can be groomed now.  First, open the Data Ports window by selecting from the upper tab: Status>Data Ports:

           [image: image20.png]V29A - Config - Data Port Status

DataPart 1
Data Packets Sent: 168581
Flood Packets Sert: 0
Packes Re Sert: 3334
Sequence Erors: 0
Packet Buffer Used: 0
Relran. Tine: 8.0
Intetface: Etheret
Facket Buffer: [N
DataFot 3
Data Packes Sent 0
Flood Packets Sert: 0
Packets Re Sert: 0
Sequence Erors: 0
Packet Bulfer Used: 1703608
Relran, Tine: 16.0
Intetface: None.

PacketButer [T

DataFort2
Data Packets Sert
Flood Packets Sent:
Packels Re Sert
Sequence Erors:
Packet Buffer Used:
Relran. Time:
Inteiface:

Packet Bufer:

DataFott 4
Data Packets Sent: 156354
Flood Packets Sert: 0
Packets Re Sert; 145
Sequence Erors: 0
Packet Bulfer Used: 2143368
Relran. Tine: 30
Intetface: Seial 2

PacketButer [N

i Cese )





26) Data Port 1 should have the Packet Buffer bar hovering around 0% if real-time data is flowing to ANF.  Shortly after the installation of a station and after the online report is sent, the station may not be sending data to ANF since that has to be set up manually and checked against the reported Q330 to prevent bogus data from flowing into ANF.

27) Next, open the Quickview window by selecting from the upper Willard tab: Status>QuickView

[image: image21.png]V29A QuickView.

Chanrels to Display

W i v the W ocha I SealeLock

I~ Vols
Chanrel 1

-
Vi ?”

Duration
@0 C 30 O B0 O 1208





Max and min values shouldn’t be too far away from 1000 and the signal should appear to wiggle.  Select a duration of 120 seconds and allow enough data to accumulate in the quickview window to see if the station appears to have the correct signature.  We don’t do any spectrum analysis at this point because it is done at DMC as a more thorough check of sensor integrity.  We can spot things such as flat-line channels, excessive noise and report them in the grooming report and to the anf mailing list.  Below is the 120s duration window (looks like a good sensor)

[image: image22.png]V29A QuickView.

Chanrels to Display
M Chl v the [ oh3 T

Chanrel 1

Duration
120 Secands of deta C0s C 30 C B0 & 2





28) Keep the QuickView window open since it will keep the Q330 from unregistering on a timeout until you are done grooming.

29) Next check in step 25 to see if Data Port 4’s packet buffer amount is around 0% or some higher number.  The new Internal balers are on all the time and so keep this buffer amount close to 0%.  The External balers allow the packet buffer amount to get up to 80% before the baler turns on and dumps all the data to the baler.  If the packet buffer amount is around 0%, indicating an internal baler, perform step 30.  Otherwise skip over step 30 to step 31.

30) For an internal baler, open a browser and go to the following URL:   station IP:5384  and hit enter.  Note: For baler44ct’s connected to a VSAT site, the port 5384 won’t connect.  Use port 5381 instead.  For the example shown below, station N26A’s IP was chosen since the example V29A we are working with does not have an internal baler, as indicted by it’s non-zero Packet Buffer in Data Port 4.  A page similar to the following will open in the browser: [image: image23.png]) PB44 Status PacketBaler44 Tag 39821 - Station TA-N26A - Mozilla Firefox
Fle Edt View Hstory Bookmarks Took Hep

@ - c a1 (1) |htip://166.241.200.168:5384/stats html

2 Most Visited

Norton- @- (@) codsatogns -

@ EarthiScope ANF Web 1M Grmail - Oriine Report "1 PBA44 Status Pack... £ | [ Facebook | Sara Pozg

Status PacketBaler44 Tag 39821 - Station TA-N26A

Baler Information

BALER44-20080825A-16K tag 39821 at 2009-01-20 22:40:55

Copyright Quanterra, Inc.
runtime: 42.991d

last baler reboot: 2008-12-08 22:53:24  reboots: S5
MEDIL GOOD

MEDIL site 2 IN USE state:
upsvolts=13.342 primaryvolts=13.700 dege=16.164
mseed record generator at 138474 and last flush at 138000

READY nedia capacity=15302.484Mb nediafres=95.939%

Q@B G [ G sl v

ST

& AT TR

75 Mapedi .05





      Note the Firmware version – Baler44-20080825A-16K in the grooming notes

31) For external balers,  check that the Packet Buffer amount is increasing in DP4 and in Willard, choose Commands>Baler Control and in the following window click on “Turn On Baler”.  You might have to click it more than once if it doesn’t grey out

    [image: image24.png]V29A - Config - Baler, Control

0330 Baler Itetface

~ Last Powerup 20030121 164303

Fonerups:8
Curent Status: Off
Time Or/Of ins

& BaiEZ
5

Tum On Baler f 0330 s o akeady requesting it to tum on

Tun OnBaler | with Web Stay-On Time (Sec) [120

Special Prevent 0330 to conect to Baler and allow Vacuum to connect

AlowVacium | for 180 seconds

Flequest Baler to Power itsef o, if 0330 indicales Baler s "Connected

Entend Web Stay-On Time sitr last Wb access-

-





Once the baler turn-on command is sent successfully, the Turn On Baler button will grey out

 [image: image25.png]V29A - Config - Baler, Control

0330 Baler Inteface
© Ethemet Last Powerup 20030121 164303

Ponerips:3
© Seral2 "

€ Gl

Curent Stalus: Manual O
Time Or/Of ins

[~Tum On Baler, i 0330 s nct aheady requesting it o un or———

Tur OniBaler | vith Web Stay n Time (3ec) [T20

[-Special Frevent 0330 o connect to Baler and allow Vacuu o connect—

AlowVacium | for 180 szconds

[-Reauest Baer to Power tsel of, 0330 indicates Baler i "Comnected—

Tun 0 Baler

[-Estend Wb Stay-On Tine aterlast Web acoess———————

Efend for [B00 Sesonds
Done. Tum On Sent





32) After a period of time the baler will turn on and data from the Packet buffer will begin to empty and some information about the buffer and DP4 tag will show up in the Willard dialog box:
[image: image26.png]Willard Q330 Setup & Control - Station V29A (166.241.200.153)
Commands  Information  Status  Configuration  Help

CaL Status ) Phose [ s Cock [IN]

Total Packets 2305 Chesksum Erors
Crd Timeauts U6 10 Enore

Out of Window 3 Average Received Bps
20090121 17:2311.000000

20090121 09:44.46:Msq From 192.168.33.14Baler! #2.20 tag 0506, TAV29A reg 20090121 16:46:25
2009.01-21 094540 Baler or DialDut Contol Set

2009.01-21 09:45:09 Meq From 192,168,331 £ Baler TA 294 readback, 122 ecs. seq end: 196804
2009.01-21 03:47.04Msq From 192.168.33.14 Baler!-2.20 tag 0506; TAV294 reg 20090121 16:45:55
2009.01-21 094838 Baler or DialDut Contol Set.

2009.01-21 0948 53Baler or DialDut Contol Set

2009.01-21 094307 Baler or DialDut Contol Set.

2003.01-21 034510 Msq From 132,168 3314 Baler TA 294 readback. 188 ecs. seq end: 200877
2009.01-21 0357:04 Baler or DialDut Contol Set

2009.01-21 0351:56:Msq From 192.168.33.14:Baler!4:2.20 tag 05095: TAV294 reg 20090121 165145





    The pertinent info is that a Baler14 (External baler) is installed at this station with       firmware version: 2.20.  The Tag on the baler data is TA-V29A, which matches the station name.  Record the firmware version on the grooming sheet.  The DP4 buffer will eventually empty and the baler will shut off.  Make sure before you disconnect from Willard that the baler has shut off and the Packet buffer for DP4 is again filling.

33) Next, check that the Configuration>Interfaces are set correctly.  Open the Configuration>Interfaces… window in Willard:
  [image: image27.png]W28A Interfaces

Seral 1| Serial 2| vl 3 Ethomet |

1P Address [182168.32.2 MAC
Options
[ Enable I~ Resticted Ping
¥ AlwaysOn [~ Resticled TCP
I” Res. QDPPing [~ Routing Enabled
I~ Unlock Always

Advanced

7 e ] X Concel | | BesePor [





With the Ethernet Tab chosen, click on “Advanced”:

[image: image28.png]W28A Baler Configuration on Ethernet  [X)
Port Type
& Telemely Only
 Telemely or Baer

X cancel| | Presenverps





34) For all new Interface Box stations and those using Ethernet, VSAT or Raven X, the above is the correct setting for Ethernet Advanced.  For a Raven station, the Ethernet Advanced should look like the following:

        [image: image29.png]113A Baler Configuration on Ethernet

Fot Type Activation
™ Recording window dane.

¥ Bulfer Fil Percent
T~ At Masimum Interval:
100 Minutes

 Telemety Oy
& Telemelry or Baler

DataPart
C1C2 036y

Baler Optons
Max Wb Bytes/s [S000

oy e
C G Pomsd P NT

Min Of [2 Feg Timeou [ Rty [13

b Sei
b 182 3]

1PAdd: [0000

Baler Ethernet
1PAdd: [192168.3314

7 o] X Concel| Detaits | @ PrservelPs





      The main thing to check is that under Activation, the Buffer Fill % is checked and set to 80%.  Make sure At maximum Interval is unchecked.

35)Cancel the advanced box and then select Serial 2:

     [image: image30.png]W28A Interfaces

Serial 1 {Serial 2| Serial 3 | Ethernet |

Paddess [6222  Baud 1152 (3]

Options

[ Enable I~ Resticted Ping
I” RTS/CTS Flow [~ Resticted TCP
I” Res. QDPPing [~ Routing Enabled
I™ Unlock Always [~ Break Detect

Masimum Byes per Second [5_ Advanced
Aoy x :am\‘ BasePort [5330





Click “Advanced”

[image: image31.png]W28A Baler on Serial 2
Pot Type Baler Actvaton
 Telmety || | Scheduled acauiston dore
 Baler ¥ Bulfer Fil Prcent [80 %

 Disivo || AtMasimum nervat

 Serial Sensor || [1440 Minutes
DataPart
1 c2 c3 G4

Baler Optons
& Power Cycled-DTR  Max Wb Bytes/s 3000

" Power Cycled -Bresk v NAPT [~ Transpaent
" Continuously Powered ¥ Commands OK.

Mi O [7 Reg. Tineou [§ ey [i3

Balr [P Addhesses

Baler Serial[152 1683314
Baler Ethermet[65.241 164135

I Urlock
X cancel| ¥ Preseive IPs





35) For all external baler installations (Interface box or not) except for those using the old-style Raven Serial modems, the above Serial 2 Advanced settings are correct.  For Interface boxes using internal balers, the Baler Option will be set to Continuously powered instead of Power Cycled – DTR as shown below:
  [image: image32.png]026A Baler on Serial 2
Fot Type Baler Actvation
€ Telemety ™ Schedled acapistion done
& Baler I~ Buffer Fil Percent [14 %
 Diakiour || AMaimum ntervak

 Serial Sensor || [100 Minutes

DataPart
c1c2 C3

Baler Optons

" Power Cycled -DTR  MaxWeb Bytes/s.

€ Power Cycled -Break [~ NAPT I Transparent
& Contiruously Powered ¥ Commands 0K

Mi O [7 Reg. Tineou [§ ey [i3

Balr [P Addhesses

Bale Seial[202100.1
Baler Ethemet[152168.33.2

Balr Config





For the internal baler settings, also the Baler IP addresses are set up differently as shown above.

36) For stations using Raven Serial modems, Serial 2 Advanced should match the following:
   [image: image33.png]113A Telemetry on Serial 2

Fot Type
& Telemety
€ Baler

€ Diakin/out

 Seral Sensor

=
X cancel| ¥ Preserve IPs





37) Next we check that the configuration tags appear correct. Click on Willard’s Configuration>Configuration Cloning …
For Raven X and Ethernet stations not using the internal baler, you should see:

[image: image34.png]W28A Configuration Cloning

SavelaFile

Configuration Base Tag [TA_Base_3ch_eth_ver06

DP1 Token Set [TA_tok_3ch_40_1_S0H_ver
DP2Token Set [disabled

DP3 Token Set [TAtok_3ch_1_SOH_ver
DP4 Token Set [TA_tok_3ch_40_1_S0H_ver

I Scide Bl Passiieid (=
HL cese & LosifonFi





For Raven Serial stations, you should see:

[image: image35.png]113A Configuration Cloning

SavelaFile

Configuration Base Tag [TA_Base_3ch_cell ver25

DP1 Token Set [TA_tok_3ch_40_1_S0H_vers
DP2Token Set [disabled

DP3 Token Set [TAtok_3ch_1_SOH_vers
DP4 Token Set [TA_tok_3ch_40_1_S0H_vers

I Scide Bl Passiieid (=
HL cese & LosifonFi





For an internal baler station, the base tag and token set names have not been set up correctly yet, but it will be a variation of the cell configuration version

38) Next, we check that the announcements are set up correctly.  To open the announcement window, click on Willard’s Configuration tab and then click on Announcements:

     [image: image36.png]W28A Announcements

PR1 | PR2 | PR3 | PR4 | PRS | PRE |
PO Receiver Infomation

PRIP Addess [13223943

Router Address [192 168331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

P, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [ Seial1 [~ Serisl2

SN Unocks
© B @ Seid1 [ Sel2

s siopAl | % carcel |





Above is for a Raven X external baler station – PR IP Address is a UCSD IP.  It is also correct for a Raven station (S/N Unlocks sends Q330 serial number out Serial 1).

For an internal baler station, Q330 interface and S/N unlocks should have Serial 2 checked since that is the route the POC’s take.  Below shows the PR1 Announcements page for an internal baler station:
   [image: image37.png]026A Announcements

PR1 | PR2 | PR3 | PR4 | PRS | PRE |
PO Receiver Infomation

PRIP Addess [13223943

Router Address [192 168331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

P, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [~ Seial1 [ Serisl2

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





The only thing that should change from PR1 to 2 to 3, 4, 5, and 6 is the PR IP address.  The following 5 pages show them for an internal baler station:

 [image: image38.png]026A Announcements

PR1 PR2 | mR3 | FR | Fas | FRs |
POC Receiver rformatin

PR IF Addess 12299443

Router Addess[1E2 68331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

P, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [~ Seial1 [ Serisl2

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





[image: image39.png]026A Announcements

PR1 | PR2 PR3 | PR4 | PRS | PRE |
PO Receiver Infomation

PRIP Addess [12595 16536

Router Address [192 168331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

e, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [~ Seial1 [ Serisl2

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





[image: image40.png]026A Announcements

Pr1 | PR2 | PR3 R4 | pRs | PRs |
POC Receiver nformatin

PR IF Addss [15820230203

Router Addess[1E2 68331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

e, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [~ Seial1 [ Serisl2

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





[image: image41.png]026A Announcements

PA1 | P2 | PR3 | PR4 PAS | R |
POC Receiver nformaton

PR IF Addess [245713228

Router Addess[1E2 68331

I Stop This Announcement
I Riandom Source Port rEEUPen | 2

e, [ |

0330 Intefaces Data Pots
P13
r2rs

¥ Ethemet [~ Seial1 [ Serisl2

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





[image: image42.png]026A Announcements

Pr1| PRz | PR3 | P4 | PRS PRS |
POC Receiver nformaton

PRIP Address

Fouter Addess

I Stop This Announcement
I Riandom Source Port "’

P, [ |

0330 Intefaces Data Pots
r1r3

I~ Ethemet [~ Sefa1 [~ Sefal2 | 1 [y

SN Unocks
© Bl [ Seid1 v Sel2

s siopAl | % carcel |





39) The analog values are checked by clicking on Willard’s Status> Misc. Analog

   [image: image43.png]W28A - Config - Miscellaneous .

Boom Postions

Cromnel 1 IRIESEN Crorec + ORI
Chornel2 IMINEZSH Crornct ORI
Chornel s IMIEGIN Crornct ORI

Anslog Postive Supply:
InputVoltage:

System Temperature:
Main Curert

Main Pawer

Antenna Curent:
Sensor A Temperature:
Sensor B Temperature:
Calbrator Timeouts:

548V
1290
14C
7ama
1.02w
1ima
Unknann
Unknann
5

i Cese )





Check that all the absolute values of the Boom positions 1,2,3 are below 45.  If they are more than that, send a centering command as follows: Willard Commands> Sensor control
    [image: image44.png]Pulse Sensor Control

Contrl Linefs) to Pulse

Sensar A Certering
Sensor A Calbrate
Sensor B Calbrate

Duration in Secands
200

T
T





Select Sensor A Centering and if the sensor to be centered is a Guralp 3T, increase the Duration in Seconds to 7.0:

   [image: image45.png]Pulse Sensor Control

Contrl Linefs]to Pulse

Sensar A Calbrate
Sensor B Calbrate

Duration in Secands
704

O = >
oot | X concel |





40) Click “Apply” and monitor the QuickView screen.  It will be obvious when the centering command is applied to the sensor because all 3 channels will go to the max or minimum values of 7M and it will take some minutes for the channels to settle down and look like seismic signals again.  At that time, the absolute values of the boom positions should all be below 20.  If several Sensor A Centering commands are sent without moving the Boom positions to acceptable levels, report that there is a sensor that won’t center.  It will probably have to be dug up and releveled.

41) If the station has a Paroscientific microbarometer installed, check that it is working by clicking Willard’s Status>Serial Sensor:

           [image: image46.png]026A - Config - Serial Se..
el | sai2]

Type: Series 1000, 6000, 3000
Pressue: 0857675 bar

Int. Temp.
Humidity

Ext. Temp.





42) Type and Pressure will be blank if the Paroscientific is not communicating

43) Finally, check that the Q330 Firmware version is correct – Willard  Information> Hardware & Software

      [image: image47.png]W28A Hardware and Software

(0330 Seral Number: 0100000EBDAESCT
AME Serial Number: 0100000B64216711
Sensor A Serial Number: 1000000000000000
Sensar B Serial Number: 1000000000000000
GAPCHPT Serial Number 4153305
GAPCHP2 Serial Nuber: 0
KMI Property Tag 2097
System Software Version: 1105
Slave Pracessar Version: 1.11
Calbrator Type: QCAL330
Calbrator Versiors 1.8
Audliary Board Type:
Auilaty Board Version:
Clock Type: Motorols M12
PLD Version: 7.9

i Cese )





    This has System Software version 1.105, which is fine except for internal baler stations, which must have version 1.110 or newer.

Once everything is checked, type “fine” on the grooming checklist file for the station, unless there were problems, which you should report on the same line.  Then exit Willard by clicking on the “Exit” button on the main Willard page.  It may be necessary to close an information window or two first to enable the Exit button.
